TAWN tests for quantitatively measuring the resolution 
and sensitivity of DSCs (version 2.1)

Results form

1. Information on the used DSC, the laboratory and the researchers

Manufacturer of the DSC:

Model of the DSC:

If appropriate, the type or model of the sensor placed in the DSC:

If appropriate, the type and model of the furnace:

Type of cooling device attached to the DSC:

If appropriate, the measuring mode used for the measurements (resolution test as well as sensitivity test):

Any other information of interest regarding the DSC:

Name of the institute / department / laboratory:

Name of the person who performed the experiments:

Name of the person who analysed the measured curves:

2. Resolution test

Crucibles:

- manufacturer, type and material:

- mass of the reference crucible:

- mass of the empty sample crucible (incl. lid):

Sample:

- is this sample supplied by the TAWN?	yes / No

- describe the sample substance as accurate as possible, i.e. supplier, purity, etc.:

- sample mass:

Purge gas:

- type:

- flow rate:

Temperature program:

- Sampling rate:

- Confirm that the temperature was exactly programmed 
as described in section 2.1 of the protocol 	Yes / No.
(i.e.	1) Isothermal at 100 °C (373 K) for 1 min;
	2) Cool from 100 °C to 25 °C (373 K - 298 K) at a cooling rate of 10 K/min;  
	3) Isothermal at 25 °C (298 K) for 1 min;
	4) Heat from 25 °C to 100 °C (298 K - 373 K) at a heating rate of 5 K/min.)
If no, indicate any differences:

- Were additional measurements performed at other heating rates? 	Yes / No. 
If yes, confirm that the temperature program was expanded with 6 additional 
series of the above mentioned 4 steps, but now with different heating rates in 
each 4th step, in the order:
10 K/min; 40 K/min; 1 K/min; 20 K/min; 3 K/min; and 5 K/min? 	Yes / No.
If no, indicate any discrepancies:

Analysis of the results:

Perform the following actions on the DSC-curve recorded during heating at a rate of 5 K/min. When other heating rates were also used, repeat all actions for each heating rate.

- Plot DSC-curves as described in section 2.2 of the protocol (i.e. a plot of heat flow versus temperature and a plot of heat flow versus time). Add the plots as appendices to this results form.

- Give the extrapolated onset temperatures (Te) of both (in most cases not completely separated) peaks. And for the double peak, give the total enthalpy of transition (trH) evaluated from the observed curve. 
	Peak	Te	trH
    	Low-temperature peak	enter value
	High-temperature peak	enter value
	Double peak		enter value
Which plot did you use for evaluating the extrapolated onset temperatures:
- heat flow versus temperature			Yes / No;
- heat flow versus time			Yes / No.
Which plot did you use for evaluating the total heat of transition:
- heat flow versus temperature			Yes / No;
- heat flow versus time			Yes / No.

- From the plot of the heat flow versus time, give the peak height and the valley value measured as described in section 2.2 of the protocol. Calculate the resolution of the DSC as the peak height divided by the valley level.
	Peak height:	enter value
	Valley level:	enter value
	Resolution:	enter value

3. Sensitivity test

Crucibles:

- manufacturer, type and material:

- mass of the reference crucible:

- mass of the empty sample crucible (incl. lid):

Sample:

- describe the sample substance as accurate as possible, i.e. supplier, purity, etc.):

- sample mass:

Purge gas:

- type:

- flow rate:

Temperature program:

- Sampling rate:

- Confirm that the temperature was exactly programmed 
as described in section 3.1 of the protocol: 	Yes / No.
(i.e.	1) Isothermal at 130 °C (403 K) for 5 min;
	2) Heat from 130 °C to 140 °C (403 K - 413 K) at a heating rate of 0.1 K/min.
If no, indicate any differences.

- Was the additional cooling curve also recorded and programmed as described? 	Yes / No. 
If yes, confirm that the temperature was programmed as described?	Yes / No.
(i.e. was the temperature program expanded with the third step:
	3) Cool from 140 °C to 130 °C (413 K - 403 K) at a cooling rate of 0.1 K/min.)
If no, indicate any discrepancies:

Analysis of the results:

- The raw data must be presented and analysed, without any data manipulation, like smoothing or deconvolution. If one wants, manipulated results can also be presented and analysed, but always in addition to the raw date. This allows to observe the effect of the data manipulation. Since the resolution and sensitivity of a DSC are not independent, perform exactly the same manipulation procedures on both the resolution test and the sensitivity test.

- Plot the DSC-curves that were recorded during the second and, if appropriate, third step of the temperature program. The horizontal axis of both plots should display the temperature, the vertical axis should display the heat flow. An endothermic and exothermic peak should be present in the heating and cooling curve, respectively. These curves are used for the evaluations. As additional information, please also plot the DSC-curves as a function of time. Add the plots as appendices to this results form.

- For each peak, give the extrapolated onset temperature (Te) and the enthalpy of transition (trH) evaluated from the observed peaks:
	Heating / Cooling	Te	trH
    	Heating	enter value	enter value
	Cooling	enter value	enter value

- Measure the top-to-top noise level and the peak height as described in section 3.2 of the protocol. Calculate the sensitivity as the peak height divided by the top-to-top noise level.

	Heating / Cooling	Top-to-top noise level	Peak height	Sensitivity
    	Heating	enter value	enter value	enter value
	Cooling	enter value	enter value	enter value